Assessors Judgment Guide

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Calibrate Electrical Equipment |
| Candidate Details | Name:_______ _______________________________________ Registration/Roll Number: ________________________ Signature: ____________________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________ Assessor’s code:___________________ Signature:__________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Perform Pre-Calibration activity - Calibrate Ammeter - Calibrate Voltmeter - Calibrate Ohmmeter - Calibrate multimeter - Calibrate 2-terminal/4-terminal Resistor(Direct method) - Calibrate 2-terminal/4-terminal Resistor(Indirect method) - Calibrate Standard Capacitor - Calibrate Standard Inductor - Prepare calibration report |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Perform functional and physical check of under test equipment/UUT (Unit Under Test) |  |  |  |  |  |
|  | Prepare equipment receipt with specifications and allot unique code. |  |  |  |  |  |
|  | Keep the equipment in controlled environment for thermal stabilization as per its manual. |  |  |  |  |  |
|  | Turn ON the equipment for warm-up as per equipment manual |  |  |  |  |  |
|  | Record the under test equipment specifications, environmental condition of laboratory. |  |  |  |  |  |
|  | Select the Reference/Standard Current Source and connecting cable on the basis of under test ammeter range. |  |  |  |  |  |
|  | Connect the under test ammeter with Reference/Standard Current Source at current terminals. |  |  |  |  |  |
|  | Select calibration points on under test ammeter such that it may cover the whole range. |  |  |  |  |  |
|  | Perform zeroing on under test (if zero function available) otherwise record zero error. |  |  |  |  |  |
|  | Verify the connections are as per connection diagram in the equipment manual before applying current |  |  |  |  |  |
|  | Apply the Reference current and record the value displayed on the under test ammeter and Reference. |  |  |  |  |  |
|  | Change the reference value and observe it on under test equipment. |  |  |  |  |  |
|  | Take at least three set of measurement on each calibration point. |  |  |  |  |  |
|  | Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |  |  |  |
|  | Stop the current source and remove the connecting cables. |  |  |  |  |  |
|  | Select the Reference/Standard Voltage Source. |  |  |  |  |  |
|  | Connect the under test voltmeter with Reference/Standard Voltage Source at voltage terminals. |  |  |  |  |  |
|  | Select calibration points on under test voltmeter such that it may cover the whole range. |  |  |  |  |  |
|  | Perform zeroing on under test (if zero function available) otherwise record zero error. |  |  |  |  |  |
|  | Verify the connections are as per connection diagram in the equipment manual before applying voltage |  |  |  |  |  |
|  | Apply the Reference voltage and record the value displayed on the under test voltmeter and Reference. |  |  |  |  |  |
|  | Change the reference value and observe it on under test equipment. |  |  |  |  |  |
|  | Take at least three set of measurement on each calibration point. |  |  |  |  |  |
|  | Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |  |  |  |
|  | Stop the voltage source and remove the connecting cables. |  |  |  |  |  |
|  | Select the Reference/Standard Resistance Source(Fixed/Discrete). |  |  |  |  |  |
|  | Connect the under test ohmmeter with Reference/Standard Resistance Source at resistance terminals. |  |  |  |  |  |
|  | Select calibration points on under test ohmmeter such that it may cover the whole range. |  |  |  |  |  |
|  | Check zero error by shortening the probes and perform zeroing on under test (if zero function/adjustment available) otherwise record zero error. |  |  |  |  |  |
|  | Verify the connections are as per connection diagram in the equipment manual before applying resistance |  |  |  |  |  |
|  | Apply the Reference resistance and record the value displayed on the under test ohmmeter and Reference. |  |  |  |  |  |
|  | Change the reference value and observe it on under test equipment. |  |  |  |  |  |
|  | Take at least three set of measurement on each calibration point. |  |  |  |  |  |
|  | Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |  |  |  |
|  | Stop the resistance source and remove the connecting cables. |  |  |  |  |  |
|  | Select the Reference/Standard according to the multimeter specifications. |  |  |  |  |  |
|  | Connect the under test multimeter with Reference/Standard as per selected parameter. |  |  |  |  |  |
|  | Select parameter(Voltage AC/DC, Current AC/DC, Resistance, Capacitance) on both Reference and under test equipment one by one. |  |  |  |  |  |
|  | Check zero error for each parameter and nullify (if zero function/adjustment available) otherwise record zero error. |  |  |  |  |  |
|  | Verify the connections are as per connection diagram in the equipment manual before applying Reference value for each parameter |  |  |  |  |  |
|  | Apply the Reference value of each parameter and record the value displayed on the under test multimeter and Reference. |  |  |  |  |  |
|  | Take at least three set of measurement on each calibration point. |  |  |  |  |  |
|  | Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |  |  |  |
|  | Stop the Standard Reference and remove the connecting cables. |  |  |  |  |  |
|  | Select the Resistance function on Reference/Standard RCL Meter |  |  |  |  |  |
|  | Short the connecting probes of RCL meter and zero the meter. |  |  |  |  |  |
|  | Connect the probes of RCL meter with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor |  |  |  |  |  |
|  | Verify the connections are as per connection diagram. |  |  |  |  |  |
|  | Note the value marked on resistor and record the value displayed on the Reference RCL meter. |  |  |  |  |  |
|  | Take at least three set of measurement. |  |  |  |  |  |
|  | Remove the connecting probes of RCL meter |  |  |  |  |  |
|  | Connect the probes of Current Source with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor |  |  |  |  |  |
|  | Select rated current to be applied to UUT and measure the corresponding value of voltage drop at the voltage terminals. |  |  |  |  |  |
|  | Put the applied values of current and measured value of voltage drop in Ohm’s Law i.e. (V = I x R) and calculate the corresponding resistance. |  |  |  |  |  |
|  | Take at least three set of measurement. |  |  |  |  |  |
|  | Remove the connecting probes of RCL meter |  |  |  |  |  |
|  | Select the Capacitance function on Reference/Standard RCL Meter |  |  |  |  |  |
|  | Short the connecting probes of RCL meter and zero the meter. |  |  |  |  |  |
|  | Connect the probes of RCL meter with Capacitor terminals |  |  |  |  |  |
|  | Verify the connections are as per connection diagram. |  |  |  |  |  |
|  | Note the value marked on Capacitor and record the value displayed on the Reference RCL meter. |  |  |  |  |  |
|  | Take at least three set of measurement. |  |  |  |  |  |
|  | Remove the connecting probes of RCL meter |  |  |  |  |  |
|  | Select the Inductance function on Reference/Standard RCL Meter |  |  |  |  |  |
|  | Short the connecting probes of RCL meter and zero the meter. |  |  |  |  |  |
|  | Connect the probes of RCL meter with Inductor terminals |  |  |  |  |  |
|  | Verify the connections are as per connection diagram. |  |  |  |  |  |
|  | Note the value marked on Inductor and record the value displayed on the Reference RCL meter. |  |  |  |  |  |
|  | Take at least three set of measurement. |  |  |  |  |  |
|  | Remove the connecting probes of RCL meter |  |  |  |  |  |
|  | Prepare data sheet by filling in the details of equipment under test and reference equipment to be used. |  |  |  |  |  |
|  | Identify and mention the procedure details and traceability statement. |  |  |  |  |  |
|  | Calculate the average, error, standard deviation and type-A uncertainty from the measured values. |  |  |  |  |  |
|  | Identify and add B-type uncertainties from the factors affecting measurement |  |  |  |  |  |
|  | Calculate the Combined and expanded uncertainty from type A and B uncertainties |  |  |  |  |  |
|  | Report the final result which includes nominal value, measured value and expanded uncertainty. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Calibrate Electrical Equipment |
| Candidate Details | Name:_______________________________________________________________________ Registration/Roll Number: __________________ Candidate Signature:__________________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor:__________________________ Assessor’s code:_________________ Signature of the Assessor:_______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Explain Calibration and technical procedure for the calibration of ammeter, voltmeter, ohmmeter, multimeter, inductor, capacitor and resistor. |  |  |
|  | Differentiate between Reference Standard and (Unit Under Test)UUT. |  |  |
|  | Explain resolution, standard deviation, uncertainty. |  |  |
|  | Explain difference between precision and accuracy |  |  |
|  | Differentiate between fixed and discrete resistor |  |  |
|  | Explain calculation method of error off error and uncertainty budget |  |  |
|  | What is the coverage factor |  |  |

Feedback to the Candidate

Candidate’s Signature________________________ Assessor’s Signature __________________

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | Electrical Instruments and Calibration |
| Competency Standards | Calibrate Electrical Equipment |
| Assessment Task | - Perform Pre-Calibration activity |
|  | - Calibrate Ammeter |
|  | - Calibrate Voltmeter |
|  | - Calibrate Ohmmeter |
|  | - Calibrate multimeter |
|  | - Calibrate 2-terminal/4-terminal Resistor(Direct method) |
|  | - Calibrate 2-terminal/4-terminal Resistor(Indirect method) |
|  | - Calibrate Standard Capacitor |
|  | - Calibrate Standard Inductor |
|  | - Prepare calibration report |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Perform functional and physical check of under test equipment/UUT (Unit Under Test) |  |  |
| - Prepare equipment receipt with specifications and allot unique code. |  |  |
| - Keep the equipment in controlled environment for thermal stabilization as per its manual. |  |  |
| - Turn ON the equipment for warm-up as per equipment manual |  |  |
| - Record the under test equipment specifications, environmental condition of laboratory. |  |  |
| - Select the Reference/Standard Current Source and connecting cable on the basis of under test ammeter range. |  |  |
| - Connect the under test ammeter with Reference/Standard Current Source at current terminals. |  |  |
| - Select calibration points on under test ammeter such that it may cover the whole range. |  |  |
| - Perform zeroing on under test (if zero function available) otherwise record zero error. |  |  |
| - Verify the connections are as per connection diagram in the equipment manual before applying current |  |  |
| - Apply the Reference current and record the value displayed on the under test ammeter and Reference. |  |  |
| - Change the reference value and observe it on under test equipment. |  |  |
| - Take at least three set of measurement on each calibration point. |  |  |
| - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |
| - Stop the current source and remove the connecting cables. |  |  |
| - Select the Reference/Standard Voltage Source. |  |  |
| - Connect the under test voltmeter with Reference/Standard Voltage Source at voltage terminals. |  |  |
| - Select calibration points on under test voltmeter such that it may cover the whole range. |  |  |
| - Perform zeroing on under test (if zero function available) otherwise record zero error. |  |  |
| - Verify the connections are as per connection diagram in the equipment manual before applying voltage |  |  |
| - Apply the Reference voltage and record the value displayed on the under test voltmeter and Reference. |  |  |
| - Change the reference value and observe it on under test equipment. |  |  |
| - Take at least three set of measurement on each calibration point. |  |  |
| - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |
| - Stop the voltage source and remove the connecting cables. |  |  |
| - Select the Reference/Standard Resistance Source(Fixed/Discrete). |  |  |
| - Connect the under test ohmmeter with Reference/Standard Resistance Source at resistance terminals. |  |  |
| - Select calibration points on under test ohmmeter such that it may cover the whole range. |  |  |
| - Check zero error by shortening the probes and perform zeroing on under test (if zero function/adjustment available) otherwise record zero error. |  |  |
| - Verify the connections are as per connection diagram in the equipment manual before applying resistance |  |  |
| - Apply the Reference resistance and record the value displayed on the under test ohmmeter and Reference. |  |  |
| - Change the reference value and observe it on under test equipment. |  |  |
| - Take at least three set of measurement on each calibration point. |  |  |
| - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |
| - Stop the resistance source and remove the connecting cables. |  |  |
| - Select the Reference/Standard according to the multimeter specifications. |  |  |
| - Connect the under test multimeter with Reference/Standard as per selected parameter. |  |  |
| - Select parameter(Voltage AC/DC, Current AC/DC, Resistance, Capacitance) on both Reference and under test equipment one by one. |  |  |
| - Check zero error for each parameter and nullify (if zero function/adjustment available) otherwise record zero error. |  |  |
| - Verify the connections are as per connection diagram in the equipment manual before applying Reference value for each parameter |  |  |
| - Apply the Reference value of each parameter and record the value displayed on the under test multimeter and Reference. |  |  |
| - Take at least three set of measurement on each calibration point. |  |  |
| - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. |  |  |
| - Stop the Standard Reference and remove the connecting cables. |  |  |
| - Select the Resistance function on Reference/Standard RCL Meter |  |  |
| - Short the connecting probes of RCL meter and zero the meter. |  |  |
| - Connect the probes of RCL meter with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor |  |  |
| - Verify the connections are as per connection diagram. |  |  |
| - Note the value marked on resistor and record the value displayed on the Reference RCL meter. |  |  |
| - Take at least three set of measurement. |  |  |
| - Remove the connecting probes of RCL meter |  |  |
| - Connect the probes of Current Source with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor |  |  |
| - Select rated current to be applied to UUT and measure the corresponding value of voltage drop at the voltage terminals. |  |  |
| - Put the applied values of current and measured value of voltage drop in Ohm’s Law i.e. (V = I x R) and calculate the corresponding resistance. |  |  |
| - Take at least three set of measurement. |  |  |
| - Remove the connecting probes of RCL meter |  |  |
| - Select the Capacitance function on Reference/Standard RCL Meter |  |  |
| - Short the connecting probes of RCL meter and zero the meter. |  |  |
| - Connect the probes of RCL meter with Capacitor terminals |  |  |
| - Verify the connections are as per connection diagram. |  |  |
| - Note the value marked on Capacitor and record the value displayed on the Reference RCL meter. |  |  |
| - Take at least three set of measurement. |  |  |
| - Remove the connecting probes of RCL meter |  |  |
| - Select the Inductance function on Reference/Standard RCL Meter |  |  |
| - Short the connecting probes of RCL meter and zero the meter. |  |  |
| - Connect the probes of RCL meter with Inductor terminals |  |  |
| - Verify the connections are as per connection diagram. |  |  |
| - Note the value marked on Inductor and record the value displayed on the Reference RCL meter. |  |  |
| - Take at least three set of measurement. |  |  |
| - Remove the connecting probes of RCL meter |  |  |
| - Prepare data sheet by filling in the details of equipment under test and reference equipment to be used. |  |  |
| - Identify and mention the procedure details and traceability statement. |  |  |
| - Calculate the average, error, standard deviation and type-A uncertainty from the measured values. |  |  |
| - Identify and add B-type uncertainties from the factors affecting measurement |  |  |
| - Calculate the Combined and expanded uncertainty from type A and B uncertainties |  |  |
| - Report the final result which includes nominal value, measured value and expanded uncertainty. |  |  |

Candidate’s Signature___________________		Assessor’s Signature_____________________

Date: ________________________________

Instruction Sheet for the Candidate

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Calibrate Electrical Equipment |
| Purpose of Assessment | Formative Assessment |

| Candidate Details | Name__________________________________________________ Registration/Roll Number_____________________________________________ |
| --- | --- |
| Guidance for Candidate | To meet this standard you are required to complete the following within the given time frame (for practical demonstration & assessment): - Perform Pre-Calibration activity - Calibrate Ammeter - Calibrate Voltmeter - Calibrate Ohmmeter - Calibrate multimeter - Calibrate 2-terminal/4-terminal Resistor(Direct method) - Calibrate 2-terminal/4-terminal Resistor(Indirect method) - Calibrate Standard Capacitor - Calibrate Standard Inductor - Prepare calibration report |
| Time: 12 hours | During a practical assessment, under observation by an assessor, you are required to - Perform Pre-Calibration activity - Perform functional and physical check of under test equipment/UUT (Unit Under Test) - Prepare equipment receipt with specifications and allot unique code. - Keep the equipment in controlled environment for thermal stabilization as per its manual. - Turn ON the equipment for warm-up as per equipment manual - Record the under test equipment specifications, environmental condition of laboratory. - Calibrate Ammeter - Select the Reference/Standard Current Source and connecting cable on the basis of under test ammeter range. - Connect the under test ammeter with Reference/Standard Current Source at current terminals. - Select calibration points on under test ammeter such that it may cover the whole range. - Perform zeroing on under test (if zero function available) otherwise record zero error. - Verify the connections are as per connection diagram in the equipment manual before applying current - Apply the Reference current and record the value displayed on the under test ammeter and Reference. - Change the reference value and observe it on under test equipment. - Take at least three set of measurement on each calibration point. - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. - Stop the current source and remove the connecting cables. - Calibrate Voltmeter - Select the Reference/Standard Voltage Source. - Connect the under test voltmeter with Reference/Standard Voltage Source at voltage terminals. - Select calibration points on under test voltmeter such that it may cover the whole range. - Perform zeroing on under test (if zero function available) otherwise record zero error. - Verify the connections are as per connection diagram in the equipment manual before applying voltage - Apply the Reference voltage and record the value displayed on the under test voltmeter and Reference. - Change the reference value and observe it on under test equipment. - Take at least three set of measurement on each calibration point. - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. - Stop the voltage source and remove the connecting cables. - Calibrate Ohmmeter - Select the Reference/Standard Resistance Source(Fixed/Discrete). - Connect the under test ohmmeter with Reference/Standard Resistance Source at resistance terminals. - Select calibration points on under test ohmmeter such that it may cover the whole range. - Check zero error by shortening the probes and perform zeroing on under test (if zero function/adjustment available) otherwise record zero error. - Verify the connections are as per connection diagram in the equipment manual before applying resistance - Apply the Reference resistance and record the value displayed on the under test ohmmeter and Reference. - Change the reference value and observe it on under test equipment. - Take at least three set of measurement on each calibration point. - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. - Stop the resistance source and remove the connecting cables. - Calibrate multimeter - Select the Reference/Standard according to the multimeter specifications. - Connect the under test multimeter with Reference/Standard as per selected parameter. - Select parameter(Voltage AC/DC, Current AC/DC, Resistance, Capacitance) on both Reference and under test equipment one by one. - Check zero error for each parameter and nullify (if zero function/adjustment available) otherwise record zero error. - Verify the connections are as per connection diagram in the equipment manual before applying Reference value for each parameter - Apply the Reference value of each parameter and record the value displayed on the under test multimeter and Reference. - Take at least three set of measurement on each calibration point. - Take values in ascending(zero to maximum) order then descending(max to zero) order and again ascending order. - Stop the Standard Reference and remove the connecting cables. - Calibrate 2-terminal/4-terminal Resistor(Direct method) - Select the Resistance function on Reference/Standard RCL Meter - Short the connecting probes of RCL meter and zero the meter. - Connect the probes of RCL meter with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor - Verify the connections are as per connection diagram. - Note the value marked on resistor and record the value displayed on the Reference RCL meter. - Take at least three set of measurement. - Remove the connecting probes of RCL meter - Calibrate 2-terminal/4-terminal Resistor(Indirect method) - Connect the probes of Current Source with the 2-terminal resistor/Current terminals(C1,C2) of 4-terminal resistor - Select rated current to be applied to UUT and measure the corresponding value of voltage drop at the voltage terminals. - Put the applied values of current and measured value of voltage drop in Ohm’s Law i.e. (V = I x R) and calculate the corresponding resistance. - Take at least three set of measurement. - Remove the connecting probes of Current source - Calibrate Standard Capacitor - Select the Capacitance function on Reference/Standard RCL Meter - Short the connecting probes of RCL meter and zero the meter. - Connect the probes of RCL meter with Capacitor terminals - Verify the connections are as per connection diagram. - Note the value marked on Capacitor and record the value displayed on the Reference RCL meter. - Take at least three set of measurement. - Remove the connecting probes of RCL meter - Calibrate Standard Inductor - Select the Inductance function on Reference/Standard RCL Meter - Short the connecting probes of RCL meter and zero the meter. - Connect the probes of RCL meter with Inductor terminals - Verify the connections are as per connection diagram. - Note the value marked on Inductor and record the value displayed on the Reference RCL meter. - Take at least three set of measurement. - Remove the connecting probes of RCL meter - Prepare calibration report - Prepare data sheet by filling in the details of equipment under test and reference equipment to be used. - Identify and mention the procedure details and traceability statement. - Calculate the average, error, standard deviation and type-A uncertainty from the measured values. - Identify and add B-type uncertainties from the factors affecting measurement - Calculate the Combined and expanded uncertainty from type A and B uncertainties - Report the final result which includes nominal value, measured value and expanded uncertainty. |
| Minimum Evidence Required |  |