Assessors Judgment Guide

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Measure Temperature, Earth Resistance and Light Intensity |
| Candidate Details | Name:_______ _______________________________________ Registration/Roll Number: ________________________ Signature: ____________________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________ Assessor’s code:___________________ Signature:__________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Measure temperature with thermocouple by voltage method - Measure the earth resistance by using earth resistance tester - Measure the light intensity with lux meter |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Picked K-type thermocouple |  |  |  |  |  |
|  | Put the Sensor tip in temperature source. |  |  |  |  |  |
|  | Connected the Multimeter on the other end of thermocouple. |  |  |  |  |  |
|  | Applied temperature by source and measure the voltage generated on multi meter |  |  |  |  |  |
|  | Recorded the readings. |  |  |  |  |  |
|  | Compared the measured value of voltage against temperature with ITS-90 table for K-type TC. |  |  |  |  |  |
|  | Shorted P1 and C1 terminals on the instrument (four terminal tester) and connected to the earth electrode under test |  |  |  |  |  |
|  | Drove C2 reference rod into the earth straight-out as far from the electrode under test as possible. |  |  |  |  |  |
|  | Drove Potential reference P2 into the earth, at a set number of points, roughly on a straight line between C1 and C2. |  |  |  |  |  |
|  | Logged the Resistance readings for each P2 point |  |  |  |  |  |
|  | Plotted Measurements on a curve of resistance vs. distance. |  |  |  |  |  |
|  | Read the Correct earth resistance from the curve for the distance that is roughly 62% of the total distance between C1 and C2. |  |  |  |  |  |
|  | Reported the result. |  |  |  |  |  |
|  | Located three or more points at same height apart from each other. |  |  |  |  |  |
|  | Turned on the Lux meter and let it stabilize. |  |  |  |  |  |
|  | Took reading on each selected point. |  |  |  |  |  |
|  | Calculated the average of measurements and reported the result. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Measure Temperature, Earth Resistance and Light Intensity |
| Candidate Details | Name:_______________________________________________________________________ Registration/Roll Number: __________________ Candidate Signature:__________________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor:__________________________ Assessor’s code:_________________ Signature of the Assessor:_______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Describe different types of thermocouple, construction and temperature ranges. |  |  |
|  | Explain the working principle of thermocouple |  |  |
|  | Explain different shapes of thermocouples and its colors. |  |  |
|  | Describe the industrial application of thermocouple and its advantages. |  |  |
|  | Differentiate between thermocouple and thermistor. |  |  |
|  | Explain thermoelectric voltage |  |  |
|  | Define working principle of thermistor. |  |  |
|  | Define types of thermistor |  |  |
|  | Define the use of thermistor for current control. |  |  |
|  | Describe different methods of earth resistance testing |  |  |
|  | Explain the procedure for selection of test points and safety considerations. |  |  |
|  | Define the operating procedure of earth resistance tester |  |  |
|  | Describe light and its intensity |  |  |
|  | Explain the working principle of lux meter. |  |  |
|  | Explain the standard ranges of light intensity for different environments |  |  |

Feedback to the Candidate

Candidate’s Signature________________________ Assessor’s Signature __________________

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | Electrical Instruments and Calibration |
| Competency Standards | Measure Temperature, Earth Resistance and Light Intensity |
| Assessment Task | - Measure temperature with thermocouple by voltage method |
|  | - Measure the earth resistance by using earth resistance tester |
|  | - Measure the light intensity with lux meter |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Pick K-type thermocouple |  |  |
| - Put the Sensor tip in temperature source. |  |  |
| - Connect the Multimeter on the other end of thermocouple. |  |  |
| - Apply temperature by source and measure the voltage generated on multi meter |  |  |
| - Record the readings. |  |  |
| - Compare the measured value of voltage against temperature with ITS-90 table for K-type TC. |  |  |
| - Short P1 and C1 terminals on the instrument (four terminal tester) and connected to the earth electrode under test |  |  |
| - Drive C2 reference rod into the earth straight-out as far from the electrode under test as possible. |  |  |
| - Potential reference P2 is then driven into the earth, at a set number of points, roughly on a straight line between C1 and C2. |  |  |
| - Resistance readings are logged for each P2 point |  |  |
| - Plot Measurements on a curve of resistance vs. distance. |  |  |
| - Correct earth resistance is read from the curve for the distance that is roughly 62% of the total distance between C1 and C2. |  |  |
| - Report the result. |  |  |
| - Locate three or more points at same height apart from each other. |  |  |
| - Turn on the Lux meter and let it stabilize. |  |  |
| - Take reading on each selected point. |  |  |
| - Calculate the average of measurements and report the result |  |  |

Candidate’s Signature___________________		Assessor’s Signature_____________________

Date: ________________________________

Instruction Sheet for the Candidate

| Qualification | Electrical Instruments and Calibration |
| --- | --- |
| Competency Standard(s) | Measure Temperature, Earth Resistance and Light Intensity |
| Purpose of Assessment | Formative Assessment |

| Candidate Details | Name__________________________________________________ Registration/Roll Number_____________________________________________ |
| --- | --- |
| Guidance for Candidate | To meet this standard you are required to complete the following within the given time frame (for practical demonstration & assessment): - Measure temperature with thermocouple by voltage method - Measure the earth resistance by using earth resistance tester - Measure the light intensity with lux meter |
| Time: 3 hours | During a practical assessment, under observation by an assessor, you are required to - Measure temperature with thermocouple by voltage method - Pick K-type thermocouple - Put the Sensor tip in temperature source. - Connect the Multimeter on the other end of thermocouple. - Apply temperature by source and measure the voltage generated on multi meter - Record the readings. - Compare the measured value of voltage against temperature with ITS-90 table for K-type TC. - Measure the earth resistance by using earth resistance tester - Short P1 and C1 terminals on the instrument (four terminal tester) and connected to the earth electrode under test - Drive C2 reference rod into the earth straight-out as far from the electrode under test as possible. - Potential reference P2 is then driven into the earth, at a set number of points, roughly on a straight line between C1 and C2. - Resistance readings are logged for each P2 point - Plot Measurements on a curve of resistance vs. distance. - Correct earth resistance is read from the curve for the distance that is roughly 62% of the total distance between C1 and C2. - Report the result. - Measure the light intensity with lux meter - Locate three or more points at same height apart from each other. - Turn on the Lux meter and let it stabilize. - Take reading on each selected point. - Calculate the average of measurements and report the result. |
| Minimum Evidence Required |  |